## Sheet1
QF-750-11 Control Report - Pinch clamps | Control Report - Pinch Clamps
Revision: A / Issue Date: 12/04/2018 | Control of Pinch Clamps
Control Report
Lot / Lot: |  |  |  | Radni nalog
Entry date for pinch clamps in final control / Datum ulaska pinc klampe u finalnu kontrolu |  |  |  | RN:
Out date for pinch clamps from final control / Datum izlaska pinc klampe iz finalne kontrole
Barcode / Serial No:
Version
barcode
Visual Control |  | Result (OK / NOK) | Note: Errors:
Solenoid cable isn’t damaged / Kabl solenoida nije oštećen
Solenoid armature isn’t stucked / Kotva solenoida se ne zaglavljuje
No sharp edges / Nema oštrih ivica
Mechanical Control |  | Result (OK / NOK) | Note: Errors:
Check if the electrode touches the head / Provera da li elektroda dodiruje celjust
Check if the solenoid armature closes to the end / Proveriti da li je maksimalni zazor izmedju kotve I tela magneta 0.5mm
Length of solenoid cable / Duzina solenoid kabla (75cm ± 5mm)
The distance between the clamps in the open position / Rastojanje izmedju čeljusti klampi u max otvorenom položaju (min 5mm)
Electrical Control |  | Result | Note: Errors:
Resistance of solenoid / Otpornost solenoida | 24V (20-27)Ω tolerance ± 5%
Resistance of solenoid / Otpornost solenoida | 15V (6.5) Ω tolerance ± 5%
Short circuit/ Detektovanje kratkog spoja
Note / Napomena:
Controlled by / Kontrolisao: | Momcilo Markovic
Concession / Ustupak:
Responsible person / Odgovorno lice: | Milovan Arsic